## C1 Thermal 6_10_26
| Water in | Water out | Screw top | Bus bar | CuCr disc | Lower outside | Upper outside | Middle inside
0A | 21.7 | 21.8 | 21.6 | 21.5 | 21.6 | 21.6 | 21.6 | 21.7
100A | 21.5 | 22.1 | 22.1 | 22 | 22.4 | 22.7 | 22.3 | 22.6
150A | 21.6 | 22.5 | 22.6 | 22.3 | 23 | 23.4 | 23.1 | 23.4
200A | 22.6 | 23.7 | 23.6 | 23.2 | 24.3 | 25 | 24.5 | 25.2
250A | 22.2 | 24 | 24.2 | 22.7 | 25 | 25.8 | 24.9 | 25.7
300A | 22.3 | 24.7 | 24.7 | 22.9 | 25.9 | 27.3 | 26 | 27.2
350A | 22.5 | 25.5 | 24.9 | 23.2 | 27 | 28.8 | 27.3 | 28.9
Based on rebuilt C1

## C2 Thermal 4_13_26
| Water in | Water out | Brass disc outside | Lower coil outside | Middle coil outside | Upper coil outside | Lower coil inside | Upper coil inside
0A | 20.8 | 20.9 | 20.5 | 20.7 | 20.8 | 20.7 | 20.9 | 20.8
100A | 20.9 | 21.2 | 21.1 | 21.3 | 21.6 | 21.3 | 21.3 | 21.5
150A | 21.9 | 22.4 | 22.7 | 23 | 23.6 | 22.8 | 22.7 | 23.2
200A | 22.6 | 23.5 | 24 | 24.5 | 25.6 | 24.2 | 24.1 | 25
250A | 22.3 | 23.8 | 25 | 25.3 | 26.9 | 25.1 | 24.9 | 26.2
300A | 22.4 | 24.5 | 26 | 26.7 | 28.9 | 26.2 | 26 | 27.8
350A | 22.6 | 25.5 | 27.3 | 28.4 | 31.4 | 27.9 | 27.5 | 29.9
Based on rebuilt C2

## Thermal 3_18_26
| Water in | Water out | Brass disc | Coil lower outside | Coil upper inside | Upper bus bar | Lower bus bar | Between bus bars
0A | 21.5 | 21.6 | 21.6 | 21.7 | 21.7 | 21.7 | 21.6 | 21.7
100A | 22.3 | 22.7 | 22.9 | 23 | 23 | 30.9 | 32.2 | 31.32
150A | 22.2 | 22.6 | 23.3 | 22.8 | 23.1 | 47.8 | 52.1 | 47.1
200A | 22.6 | 23.3 | 24.5 | 23.9 | 23.9 | 61.4 | 70.6 | 65.9
Based on old top coil with 1 GPM of flow.

## Thermal 3_05_26
Current | Water in | Water out | Bottom plate | Coil bottom outside | Coil top inside | Top bus bar | Bottom bus bar | Current carrying screw
0A | 21.1 | 19.9 | 20.7 | 20.7 | 20.8 | 21 | 20.9 | 20.8
100A | 21.5 | 21.4 | 22.3 | 22.9 | 22.5 | 21.7 | 21.3 | 21.2
150A | 21.8 | 22.3 | 24.3 | 25.7 | 24.4 | 22.1 | 21.9 | 22.3
200A | 22.6 | 23.9 | 27.8 | 30.1 | 27.2 | 23 | 23 | 24
250A | 23.4 | 25.3 | 31.7 | 35.3 | 30.1 | 23.8 | 24 | 25.7
300A | 22.8 | 25.5 | 35.2 | 40.1 | 31.7 | 23.3 | 23.4 | 26.1
Based on coil 1 from 2025 with new water manifold, with 0.95 GPM in main coil, 0.59 GPM in bus bars, and 6 PSI dropped across main coil

## Coil 2 Resistance 6_15_26
C2
Current: 9.95 A | Voltage: 0.089V
Position | Voltage(mV) | dV | R (mOhm)
Handle | 0.106 | 0.106 | 0.010653
Bottom | 1.414 | 1.308 | 0.13146
1 | 1.981 | 0.567 | 0.056985
2 | 4.18 | 2.199 | 0.22101
3 | 6.512 | 2.332 | 0.23437
4 | 8.753 | 2.241 | 0.22523
5 | 11.068 | 2.315 | 0.23266
6 | 13.351 | 2.283 | 0.22945
7 | 15.64 | 2.289 | 0.23005
8 | 17.942 | 2.302 | 0.23136
9 | 20.117 | 2.175 | 0.21859
10 | 22.384 | 2.267 | 0.22784
11 | 24.653 | 2.269 | 0.22804
12 | 26.863 | 2.21 | 0.22211
13 | 29.27 | 2.407 | 0.24191
14 | 31.605 | 2.335 | 0.23467
15 | 34.004 | 2.399 | 0.24111
16 | 36.259 | 2.255 | 0.22663
17 | 38.491 | 2.232 | 0.22432
18 | 40.873 | 2.382 | 0.2394
19 | 43.121 | 2.248 | 0.22593
20 | 45.332 | 2.211 | 0.22221
21 | 47.687 | 2.355 | 0.23668
22 | 49.982 | 2.295 | 0.23065
23 | 52.289 | 2.307 | 0.23186
24 | 54.522 | 2.233 | 0.22442
25 | 56.796 | 2.274 | 0.22854
26 | 59.041 | 2.245 | 0.22563
27 | 61.279 | 2.238 | 0.22492
28 | 63.514 | 2.235 | 0.22462
29 | 66.056 | 2.542 | 0.25548
30 | 68.237 | 2.181 | 0.2192
31 | 70.449 | 2.212 | 0.22231
32 | 72.763 | 2.314 | 0.23256
Handle | 73.266 | 0.503 | 0.050553
End | 73.401 | 0.135 | 0.013568

## Coil 1 Resistance 6_15_26
C1
Current: 9.95 A | Voltage: 0.084 V
Position | Voltage(mV) | dV | R (mOhm)
Handle | 0.113 | 0.113 | 0.011357
Bottom | 1.326 | 1.213 | 0.12191
1 | 2.009 | 0.683 | 0.068643
2 | 4.33 | 2.321 | 0.23327
3 | 6.632 | 2.302 | 0.23136
4 | 8.883 | 2.251 | 0.22623
5 | 11.175 | 2.292 | 0.23035
6 | 13.43 | 2.255 | 0.22663
7 | 15.733 | 2.303 | 0.23146
8 | 18.075 | 2.342 | 0.23538
9 | 20.987 | 2.912 | 0.29266
10 | 23.234 | 2.247 | 0.22583
11 | 25.561 | 2.327 | 0.23387
12 | 27.867 | 2.306 | 0.23176
13 | 30.168 | 2.301 | 0.23126
14 | 32.44 | 2.272 | 0.22834
15 | 34.713 | 2.273 | 0.22844
16 | 37.015 | 2.302 | 0.23136
17 | 39.376 | 2.361 | 0.23729
18 | 41.758 | 2.382 | 0.2394
19 | 44.147 | 2.389 | 0.2401
20 | 46.453 | 2.306 | 0.23176
21 | 48.742 | 2.289 | 0.23005
22 | 51.067 | 2.325 | 0.23367
23 | 53.253 | 2.186 | 0.2197
24 | 55.475 | 2.222 | 0.22332
25 | 57.909 | 2.434 | 0.24462
26 | 60.282 | 2.373 | 0.23849
27 | 62.454 | 2.172 | 0.21829
28 | 64.675 | 2.221 | 0.22322
29 | 67.136 | 2.461 | 0.24734
30 | 69.327 | 2.191 | 0.2202
31 | 71.678 | 2.351 | 0.23628
32 | 73.949 | 2.271 | 0.22824
Handle | 74.442 | 0.493 | 0.049548
End | 74.55 | 0.108 | 0.010854

## Resistance 2_19_26
| C1 |  | C2 |  | Summary
Position | Voltage (mV) | Resistance (mΩ) | Voltage (mV) | Resistance (mΩ) | Position | C1 | C2
input | 0.464 | 0.15467 | 0.037 | 0.012333 | 0 | 0.14367 | 0.142
handle (in) | 0.495 | 0.010333 | 0.058 | 0.007 | 1 | 0.070333 | 0.12133
0 | 0.926 | 0.14367 | 0.484 | 0.142 | 2 | 0.24133 | 0.33233
1 | 1.137 | 0.070333 | 0.848 | 0.12133 | 3 | 0.257 | 0.27867
2 | 1.861 | 0.24133 | 1.845 | 0.33233 | 4 | 0.21433 | 0.27
3 | 2.632 | 0.257 | 2.681 | 0.27867 | 5 | 0.22833 | 0.26667
4 | 3.275 | 0.21433 | 3.491 | 0.27 | 6 | 0.249 | 0.453
5 | 3.96 | 0.22833 | 4.291 | 0.26667 | 7 | 0.24033 | 0.258
6 | 4.707 | 0.249 | 5.65 | 0.453 | 8 | 0.22933 | 0.26267
7 | 5.428 | 0.24033 | 6.424 | 0.258 | 9 | 0.23767 | 0.27767
8 | 6.116 | 0.22933 | 7.212 | 0.26267 | 10 | 0.24233 | 0.27467
9 | 6.829 | 0.23767 | 8.045 | 0.27767 | 11 | 0.26833 | 0.274
10 | 7.556 | 0.24233 | 8.869 | 0.27467 | 12 | 0.26733 | 0.34267
11 | 8.361 | 0.26833 | 9.691 | 0.274 | 13 | 0.34833 | 0.28
12 | 9.163 | 0.26733 | 10.719 | 0.34267 | 14 | 0.254 | 0.31433
13 | 10.208 | 0.34833 | 11.559 | 0.28 | 15 | 0.251 | 0.27033
14 | 10.97 | 0.254 | 12.502 | 0.31433 | 16 | 0.23567 | 0.26333
15 | 11.723 | 0.251 | 13.313 | 0.27033 | 17 | 0.256 | 0.30233
16 | 12.43 | 0.23567 | 14.103 | 0.26333 | 18 | 0.21667 | 0.25733
17 | 13.198 | 0.256 | 15.01 | 0.30233 | 19 | 0.27267 | 0.297
18 | 13.848 | 0.21667 | 15.782 | 0.25733 | 20 | 0.24833 | 0.338
19 | 14.666 | 0.27267 | 16.673 | 0.297 | 21 | 0.264 | 0.66133
20 | 15.411 | 0.24833 | 17.687 | 0.338 | 22 | 0.28267 | 0.28133
21 | 16.203 | 0.264 | 19.671 | 0.66133 | 23 | 0.28133 | 0.27067
22 | 17.051 | 0.28267 | 20.515 | 0.28133 | 24 | 0.26367 | 0.33967
23 | 17.895 | 0.28133 | 21.327 | 0.27067 | 25 | 0.32067 | 0.47
24 | 18.686 | 0.26367 | 22.346 | 0.33967 | 26 | 0.41433 | 0.28567
25 | 19.648 | 0.32067 | 23.756 | 0.47 | 27 | 0.22133 | 0.438
26 | 20.891 | 0.41433 | 24.613 | 0.28567 | 28 | 0.227 | 0.266
27 | 21.555 | 0.22133 | 25.927 | 0.438 | 29 | 0.26433 | 0.241
28 | 22.236 | 0.227 | 26.725 | 0.266 | 30 | 0.25733 | 0.242
29 | 23.029 | 0.26433 | 27.448 | 0.241 | 31 | 0.28033 | 0.25867
30 | 23.801 | 0.25733 | 28.174 | 0.242
31 | 24.642 | 0.28033 | 28.95 | 0.25867
handle (out) | 25.5 | 0.286 | 29.837 | 0.29567
output | 25.5535 | 0.017833 | 29.873 | 0.012
total |  | 8.3632 |  | 9.9453

## Coil 1 Resistance 7_29_25 (UTD)
| Total Resistance (mΩ) | Layer Resistance (mΩ)
Layer
BB1 | 0 | 0
BB1-BS | 0.036 | 0.036
BS-BP | 0.121 | 0.085
BP-1 | 0.22 | 0.099
2025-01-02 | 0.44 | 0.22
2025-02-03 | 0.665 | 0.225
2025-03-04 | 0.894 | 0.229
2025-04-05 | 1.126 | 0.232
2025-05-06 | 1.396 | 0.27
2025-06-07 | 1.647 | 0.251
2025-07-08 | 1.843 | 0.196
2025-08-09 | 2.078 | 0.235
2025-09-10 | 2.322 | 0.244
2025-10-11 | 2.559 | 0.237
2025-11-12 | 2.807 | 0.248
2025-12-13 | 3.049 | 0.242
13-14 | 3.289 | 0.24
14-15 | 3.516 | 0.227
15-16 | 3.717 | 0.201
16-17 | 3.974 | 0.257
17-18 | 4.206 | 0.232
18-19 | 4.462 | 0.256
19-20 | 4.737 | 0.275
20-21 | 5.079 | 0.342
21-22 | 5.334 | 0.255
22-23 | 5.573 | 0.239
23-24 | 5.813 | 0.24
24-25 | 6.029 | 0.216
25-26 | 6.31 | 0.281
26-27 | 6.61 | 0.3
27-28 | 6.879 | 0.269
28-29 | 7.162 | 0.283
29-30 | 7.456 | 0.294
30-31 | 7.715 | 0.259
31-BB2 | 7.729 | 0.014
Important Coil Layer Resistance Statistics (for (1-2)-(30-31))
|  |  |  | Coil 1 | Coil 2 | Both Coils
|  |  | Minimum | 0.196 | 0.191 | 0.191
|  |  | Average | 0.24983 | 0.23567 | 0.24275
|  |  | Median | 0.246 | 0.2325 | 0.23925
|  |  | Standard Deviation | 0.054338 | 0.020273 | 0.037305
|  |  | Maximum | 0.342 | 0.285 | 0.342

## Coil 2 Resistance 8_01_25 (UTD)
Layer | Total Resistance (mΩ) | Layer Resistance (m Ω)
BB1 | 0
BB1-BS | 0.036 | 0.036
BS-BP | 0.157 | 0.121
BP-1 | 0.261 | 0.104
2025-01-02 | 0.515 | 0.254
2025-02-03 | 0.759 | 0.244
2025-03-04 | 0.991 | 0.232
2025-04-05 | 1.249 | 0.258
2025-05-06 | 1.477 | 0.228
2025-06-07 | 1.694 | 0.217
2025-07-08 | 1.914 | 0.22
2025-08-09 | 2.117 | 0.203
2025-09-10 | 2.322 | 0.205
2025-10-11 | 2.549 | 0.227
2025-11-12 | 2.782 | 0.233
2025-12-13 | 3.033 | 0.251
13-14 | 3.261 | 0.228
14-15 | 3.494 | 0.233
15-16 | 3.721 | 0.227
16-17 | 3.912 | 0.191
17-18 | 4.141 | 0.229
18-19 | 4.383 | 0.242
19-20 | 4.617 | 0.234
20-21 | 4.849 | 0.232
21-22 | 5.091 | 0.242
22-23 | 5.303 | 0.212
23-24 | 5.535 | 0.232
24-25 | 5.757 | 0.222
25-26 | 6.042 | 0.285
26-27 | 6.302 | 0.26
27-28 | 6.551 | 0.249
28-29 | 6.81 | 0.259
29-30 | 7.072 | 0.262
30-31 | 7.331 | 0.259
31-BB2 | 7.479 | 0.148
|  |  | Important Coil Layer Resistance Statistics (for (1-2)-(30-31)) | Coil 1 | Coil 2 | Both Coils
|  |  | Minimum | 0.196 | 0.191 | 0.191
|  |  | Average | 0.24983 | 0.23567 | 0.24275
|  |  | Median | 0.246 | 0.2325 | 0.23925
|  |  | Standard Deviation | 0.054338 | 0.020273 | 0.037305
|  |  | Maximum | 0.342 | 0.285 | 0.342

## Coil 1 B-Field 8_04_25 (UTD)
Distance from Coil Bottom (mm) | Magnetic Field (Gauss)
-74.8538 | 126.29
-71.6788 | 135.67
-68.5038 | 146.96
-65.3288 | 163.47
-62.1538 | 184.83
-58.9788 | 200.84
-55.8038 | 229.51
-52.6288 | 249.33
-49.4538 | 269.53
-46.2788 | 296.01
-43.1038 | 323.04
-39.9288 | 346.7
-36.7538 | 367.53
-35.1663 | 375.06
-33.5788 | 383.87
-31.9913 | 390.44
-30.4038 | 394.05
-28.8163 | 397.55
-27.2288 | 404.8
-25.6413 | 406.91
-24.0538 | 408.93
-22.4663 | 410.04
-20.8788 | 407.14
-19.2913 | 406.01
-17.7038 | 403.18
-16.1163 | 396.48
-14.5288 | 392.8
-12.9413 | 380.07
-11.3538 | 372.77
-9.7663 | 362.85
-8.1788 | 349.47
-6.5913 | 337.49
-5.0038 | 327.16 | Coil Bottom Approximately Here: 2.947 inches from the top
-3.4163 | 311.56
-1.8288 | 295.7
-0.2413 | 285.49
1.3462 | 278.37
2.9337 | 256.37
4.5212 | 247.98
6.1087 | 238.64
7.6962 | 222.79
9.2837 | 213.2
10.8712 | 203.95 |  |  | Atoms approximately here 17 mm from coil bottom
12.4587 | 186.77
14.0462 | 178.21
15.6337 | 169.18
17.2212 | 161.04
18.8087 | 155.28
20.3962 | 145.65
21.9837 | 136.9
23.5712 | 116.13
25.1587 | 110.88
26.7462 | 103.46
33.0962 | 92.88
39.4462 | 71.55
45.7962 | 55.47
52.1462 | 42.79
58.4962 | 35.77
64.8462 | 30.08
71.1962 | 26.48
77.5462 | 22.1
90.2462 | 16.94
102.9462 | 12.93 |  | Important Fields |  | Measured for Coil 1 | Measured for Coil 2 | Simulated
115.6462 | 11.25 |  | Top of Plastic Piece (74.85 mm above bottom)
128.3462 | 9.7 |  |  |  | 126.29 G | 119.2 G | 118.63
|  |  | Max Field (22.4663 mm above Coil Bottom) |  | 410.04 G | 413.44 G (at 25.6413 mm above Coil Bottom) | 390.28 | (Sim max is 391.87 G at 25.167 mm above the coil bottom)
|  |  | Bottom of Coil (technically 0.2413 mm above) |  | 285.49 G | 285.69 G | 273.82
|  |  | Atom Location (atoms actually at 17 mm, but measurement location at 17.2212 mm below Coil Bottom) |  | 161.04 G | 164.34 G | 156.3

## Coil 2 B-Field 8_13_25 (UTD)
Distance from Coil Bottom (mm) | Magnetic Field (Gauss)
-79.6163 | 98.79
-78.0288 | 104.86
-76.4413 | 112.4
-74.8538 | 119.2
-73.2663 | 127.44
-71.6788 | 132.88
-70.0913 | 144.48
-68.5038 | 154.14
-66.9163 | 163.87
-65.3288 | 172.25
-63.7413 | 181.39
-62.1538 | 190.27
-60.5663 | 199.76
-58.9788 | 209.41
-57.3913 | 231.55
-55.8038 | 239.2
-54.2163 | 255.77
-52.6288 | 267.24
-51.0413 | 283.88
-49.4538 | 305.75
-47.8663 | 311.62
-46.2788 | 325.51
-44.6913 | 342.95
-43.1038 | 350.83
-41.5163 | 361.31
-39.9288 | 370.16
-38.3413 | 380.05
-36.7538 | 384.39
-35.1663 | 390.41
-33.5788 | 398.06
-31.9913 | 398.46
-30.4038 | 404.21
-28.8163 | 408.35
-27.2288 | 412.55
-25.6413 | 413.44
-24.0538 | 409.41
-22.4663 | 407.67
-20.8788 | 402.25
-19.2913 | 400.35
-17.7038 | 397.63
-16.1163 | 389.53
-14.5288 | 381.38
-12.9413 | 371.62
-11.3538 | 367.72
-9.7663 | 352.06
-8.1788 | 348.06
-6.5913 | 336.82
-5.0038 | 328.29
-3.4163 | 299.92
-1.8288 | 290.24
-0.2413 | 285.69
1.3462 | 275.77
2.9337 | 263.71
4.5212 | 256.64
6.1087 | 242.87
7.6962 | 233.66
9.2837 | 212.33
10.8712 | 200.75
12.4587 | 188.22
14.0462 | 181.13
15.6337 | 171.22
17.2212 | 164.34
18.8087 | 147.48
20.3962 | 141.58
21.9837 | 133.43
23.5712 | 125.95
25.1587 | 120.67
26.7462 | 117.34
28.3337 | 110.3
29.9212 | 104.34
31.5087 | 98.81
33.0962 | 92.84
39.4462 | 78.39
45.7962 | 59.61
52.1462 | 50.43
58.4962 | 39.94
64.8462 | 33.05
77.5462 | 28.15
90.2462 | 20.25
102.9462 | 15.47
115.6462 | 11.44
128.3462 | 9.37
|  | Important Fields | Measured for Coil 2 | Measured for Coil 1 | Simulated
|  | Top of Plastic Piece (74.85 mm above bottom) | 119.2 G | 126.29 G | 118.63
410.04 G (at 22.4663 mm above Coil Bottom)
|  | Max Field (25.6413 mm above Coil Bottom for Coil 2) | 413.44 |  | 391.8 | (Sim max is 391.87 G at 25.167 mm above the coil bottom)
|  | Bottom of Coil (technically 0.2413 mm above) | 285.69 | 285.49 | 273.82
Atom Location (atoms actually at 17 mm, but measurement location at 17.2212 mm below Coil Bottom)
|  |  | 164.34 | 161.04 | 156.3

## Coil 1 Temps 8_7_25 PC (UTD)
Current | Voltage | Power | Inlet tube | BB1 | BB2 (top, non water cooled part) | Screw | BP | Coil Low Outer (Tape layer to not ruin thermocouple BP measurement) | Coil High Inner | Outlet tube | Notes
0 | 0 | 0 | 18.6 | 17.5 | 17.9 | 17.8 | 16.9 | 17 | 17.2 | 20 | No Parallel Cooling
0 | 0 | 0 | 16.5 | 15.4 | 16.3 | 17.5 | 15 | 17.7 | 15.6 | 17.3 | Cooling On
50.11 | 0.401 | 19.4076 | 17.2 | 15.7 | 16.6 | 18 | 16 | 17.8 | 15.8 | 17
100.01 | 0.805 | 77.3055 | 17.4 | 15.9 | 17.4 | 18.3 | 16.9 | 18.6 | 16.1 | 17.4
150.03 | 1.217 | 173.9721 | 17.8 | 16 | 17.9 | 18.7 | 18.4 | 19 | 16.5 | 17.5
200.04 | 1.63 | 309.2837 | 17.6 | 16.2 | 18.7 | 18.9 | 20.6 | 20.5 | 17.3 | 18
250.04 | 2.029 | 483.2171 | 17.9 | 16.5 | 19.6 | 19.4 | 23.3 | 22.9 | 18.4 | 18.5
275.07 | 2.272 | 584.8032 | 17.7 | 16.2 | 20.1 | 19.7 | 24.4 | 23.6 | 18.5 | 18.5
300.08 | 2.489 | 695.981 | 18.1 | 16.5 | 20.7 | 20 | 26.4 | 24.8 | 19.3 | 18.9
Total Temperature Change (°C) |  |  | 1.6 | 1.1 | 4.4 | 2.5 | 11.4 | 7.1 | 3.7 | 1.6
Experiment Details
Cooling System: Parallel
Water Flow: 1.75 gal/min
Flow to Coil: 1.28 gal/min
Flow to Bus Bars 0.47 gal/min
Inward Pressure: 25 PSI
Outward Pressure: 10 PSI
Cooling Source: Chiller set to 14°C
Time Between Tempeatures: 6 minutes

## Coil 2 Temps 8_7_25 PC (UTD)
Current | Voltage | Power | Inlet tube | BB1 | BB2 (top, non water cooled part) | Screw | BP | Coil Low Outer (Tape layer to not ruin thermocouple BP measurement) | Coil High Inner | Outlet tube | Notes
0 | 0 | 0 | 21.9 | 21.6 | 21.6 | 21.6 | 21.2 | 21.4 | 21.5 | 22.1 | No Cooling
0 | 0 | 0 | 16.2 | 15.1 | 16.2 | 16.4 | 14.5 | 16 | 15.6 | 17.6 | Cooling On
50.11 | 0.384 | 18.7799 | 17.9 | 17.8 | 17.4 | 17.9 | 16.1 | 17.7 | 16.7 | 18.1
100.01 | 0.774 | 74.805 | 18 | 17.9 | 18.3 | 18.3 | 16.9 | 18.5 | 17.2 | 18.4
150.03 | 1.169 | 168.3448 | 18 | 18 | 19.8 | 18.9 | 18.2 | 19.4 | 17.8 | 18.5
200.04 | 1.538 | 299.2797 | 18.1 | 18.3 | 21.7 | 19.8 | 20.1 | 21.2 | 19.2 | 19
250.06 | 1.981 | 467.6619 | 18.3 | 18.5 | 24.3 | 21.2 | 22.5 | 23.3 | 20.6 | 19.5
275.07 | 2.176 | 565.8874 | 18.2 | 18.6 | 24.2 | 21.7 | 23.5 | 24.4 | 21.3 | 19.4
300.08 | 2.376 | 673.469 | 18.4 | 18.9 | 25.7 | 22.5 | 25.2 | 26 | 22.4 | 19.9
330 | 2.697 | 814.4631 | 18.6 | 19.1 | 27.6 | 23.4 | 27.2 | 28.2 | 23.9 | 20.2
Total Temperature Change (°C) |  |  | 2.4 | 4 | 11.4 | 7 | 12.7 | 12.2 | 8.3 | 2.6
Experiment Details
Cooling System: Parallel
Water Flow: 1.75 gal/min
Flow to Coil: 1.28 gal/min
Flow to Bus Bars 0.47 gal/min
Inward Pressure: 25 PSI
Outward Pressure: 10 PSI
Cooling Source: Chiller set to 14°C
Time Between Tempeatures: 6 minutes

## Bitter Coil Impedance 1_25_24
Frequency (Hz) | Impedence (Ω) | Coil Voltage (mv) | 100 Ohm Resistor Voltage (mv) | Angular Frequency (rad/s)
2 | 0.007966101695 | 18.8 | 236 | 12.56637061
5 | 0.008333333333 | 20.0 | 240 | 31.41592654
10 | 0.008253968254 | 20.8 | 252 | 62.83185307
20 | 0.009836065574 | 24.0 | 244 | 125.6637061
40 | 0.01466666667 | 35.2 | 240 | 251.3274123
80 | 0.026 | 62.4 | 240 | 502.6548246
160 | 0.04666666667 | 112 | 240 | 1005.309649
320 | 0.09333333333 | 224 | 240 | 2010.619298
640 | 0.1770491803 | 432 | 244 | 4021.238597
1000 | 0.2591666667 | 622 | 240 | 6283.185307
2000 | 0.55 | 1320 | 240 | 12566.37061
4000 | 1.066666667 | 2560 | 240 | 25132.74123
8000 | 2.0625 | 5280 | 256 | 50265.48246
10000 | 2.545454545 | 6720 | 264 | 62831.85307
20000 | 4.609375 | 11800 | 256 | 125663.7061
Fitted Inductance: |  | 42.6 ±1.0 µH
Fitted Resistance: |  | 8.1 ±0.3 mΩ

## Coil Temps SC 7_23_25
Current | Voltage | Inlet tube | BB1 | BB2 | Screw | BP | Coil Low Outer | Coil High Inner | Outlet tube
0 | 0 | 17.7 | 16.5 | 16.5 | 17.8 | 16.6 | 16.5 | 16.4 | 17.4
49.99 | 0.484 | 18.4 | 16.8 | 16.9 | 18.9 | 17.2 | 17.4 | 16.8 | 18.5
100.01 | 0.9 | 17.6 | 16.8 | 16.9 | 17.7 | 18.8 | 17.5 | 17.4 | 18.9
150.03 | 1.3 | 18.8 | 17.1 | 17.3 | 18.4 | 21 | 18.8 | 18.6 | 19.1
200.04 | 1.725 | 17.8 | 17.4 | 17.3 | 18.7 | 24 | 20.4 | 20 | 19.1
250.06 | 2.142 | 18.1 | 17.5 | 17.4 | 19.6 | 27.9 | 22.5 | 22 | 19.4
275.07 | 2.372 | 18.1 | 17.6 | 17.4 | 19.9 | 30 | 23.9 | 23.5 | 18.4
300.08 | 2.593 | 19.1 | 19.2 | 18.3 | 20.7 | 32.5 | 28.7 | 25.2 | 19.7
330.02 | 2.867 | 19.3 | 19.6 | 18.5 | 21.3 | 35.6 | 31.1 | 27.2 | 19.8
Temp Difference (°C) |  | 1.6 | 3.1 | 2 | 3.5 | 19 | 14.6 | 10.8 | 2.4
Water Flow |  | 0.73 gal/min to coil and bus bars
Water Pressure |  | 38 PSI
Cooling Type: Series to Coil then Bus Bars from Chiller.
Chiller Temperature: |  | 15°C

## Coil 1 Temps 7_30_25 SC
Current (A) | Power (W) | Voltage (V) | Inlet tube | BB1 | BB2 (top, non water cooled part) | Screw | BP | Coil Low Outer | Coil High Inner | Outlet tube
0 | 0 | 0 | 21 | 20.9 | 20.9 | 21 | 20.9 | 20.9 | 21 | 21.8 | No Cooling
0 | 0 | 0 | 15.9 | 15 | 15.5 | 15.1 | 15.4 | 15.6 | 14.9 | 16.4 | Cooling On
50.11 | 18.3053 | 0.484 | 16.4 | 15.9 | 17 | 16.2 | 16.5 | 17 | 15.5 | 17.6
100.01 | 72.9146 | 0.892 | 16.5 | 16.3 | 18.9 | 16.6 | 17.7 | 18.1 | 16 | 17.7
150.03 | 164.0906 | 1.313 | 16.6 | 16.8 | 21 | 17.1 | 19.8 | 19.5 | 16.7 | 18.1
200.04 | 291.7167 | 1.721 | 16.8 | 17.6 | 23.9 | 18 | 22.6 | 21.7 | 17.8 | 18.5
250.06 | 455.8437 | 2.146 | 17 | 18.4 | 27 | 19.1 | 26.4 | 24.7 | 19.3 | 19.2
275.07 | 551.587 | 2.363 | 17.2 | 18.7 | 28.2 | 19.5 | 28.2 | 26.2 | 19.5 | 19.2
300.08 | 656.45 | 2.571 | 17.1 | 19.3 | 29.8 | 20.3 | 30.8 | 28.5 | 20.9 | 19.9
Temp Difference (°C) |  |  | 1.2 | 4.3 | 14.3 | 5.2 | 15.4 | 12.9 | 6 | 3.5
Water Flow |  | 0.73 gal/min to coil and bus bars
Water Pressure |  | 38 PSI
Cooling Type: Series to Coil then Bus Bars from Chiller.
Chiller Temperature: |  | 15°C

## Coil 1 Resistance 7_28_25
Layer | Resistance | Layer Resistance
BB1 | 0 | 0
BS | 0.072 | 0.072
BP | 0.196 | 0.124
1 | 0.336 | 0.14
2 | 0.568 | 0.232
3 | 0.825 | 0.257
4 | 1.18 | 0.355
5 | 1.35 | 0.17
6 | 1.892 | 0.542
7 | 2.126 | 0.234
8 | 2.375 | 0.249
9 | 2.589 | 0.214
10 | 2.845 | 0.256
11 | 3.132 | 0.287
12 | 3.462 | 0.33
13 | 3.748 | 0.286
14 | 4.021 | 0.273
15 | 4.277 | 0.256
16 | 4.498 | 0.221
17 | 4.737 | 0.239
18 | 4.998 | 0.261
19 | 5.291 | 0.293
20 | 5.624 | 0.333
21 | 6.12 | 0.496
22 | 6.406 | 0.286
23 | 6.645 | 0.239
24 | 6.917 | 0.272
25 | 7.226 | 0.309
26 | 7.512 | 0.286
27 | 7.812 | 0.3
28 | 8.066 | 0.254
29 | 8.466 | 0.4
30 | 8.74 | 0.274
31 | 8.995 | 0.255
BB2 | 9.048 | 0.053

## Old Coil 1 Resistance 7_21_25
Layer | Total Resistance (mOhms) | Resistance (mOhms)
BB1 | 0
BP | 1.027 | 1.027
1 | 1.155 | 0.128 | Average | 0.19594
2 | 1.268 | 0.113 | Median | 0.208
3 | 1.513 | 0.245 | Max | 0.265
4 | 1.749 | 0.236 | Min | 0.087
5 | 2.014 | 0.265 | 3rd Quartile | 0.237
6 | 2.231 | 0.217 | 1st Quartile | 0.1705
7 | 2.447 | 0.216
8 | 2.66 | 0.213
9 | 2.896 | 0.236
10 | 3.11 | 0.214
11 | 3.35 | 0.24 | Top-Bottom Resistance | 6.074
12 | 3.552 | 0.202
13 | 3.712 | 0.16
14 | 3.947 | 0.235
15 | 4.185 | 0.238
16 | 4.38 | 0.195
17 | 4.561 | 0.181
18 | 4.765 | 0.204
19 | 5.018 | 0.253
20 | 5.146 | 0.128
21 | 5.392 | 0.246
22 | 5.532 | 0.14
23 | 5.619 | 0.087
24 | 5.865 | 0.246
25 | 6.108 | 0.243
26 | 6.23 | 0.122
27 | 6.363 | 0.133
28 | 6.554 | 0.191
29 | 6.747 | 0.193
30 | 6.955 | 0.208
31 | 7.101 | 0.146
BB2 | 8.126 | 1.025

## Old Coil 2 Resistance 7_21_25
Layer | Total Resistance (mΩ) | Resistance (mΩ)
BB2 | 8.748
BP | 7.89 | 0.858
1 | 7.65 | 0.24
2 | 7.431 | 0.219
3 | 7.144 | 0.287
4 | 6.944 | 0.2
5 | 6.724 | 0.22
6 | 6.536 | 0.188
7 | 6.366 | 0.17
8 | 5.994 | 0.372
9 | 5.737 | 0.257
10 | 5.479 | 0.258
11 | 5.363 | 0.116
12 | 5.1 | 0.263
13 | 4.731 | 0.369
14 | 4.458 | 0.273
15 | 4.181 | 0.277
16 | 3.9 | 0.281
17 | 3.621 | 0.279
18 | 3.352 | 0.269
19 | 3.061 | 0.291
20 | 2.737 | 0.324
21 | 2.446 | 0.291
22 | 2.175 | 0.271
23 | 1.907 | 0.268
24 | 1.771 | 0.136
25 | 1.467 | 0.304
26 | 1.208 | 0.259
27 | 1.064 | 0.144
28 | 0.801 | 0.263
29 | 0.553 | 0.248
30 | 0.351 | 0.202
31 | 0.102 | 0.249
BB2 | 0 | 0.102

## Coil 1 Resistance 7_8_25
Layer | Layer Resistance (mΩ) | Total Resistance (mΩ)
BB1
BB1-BS | 0.053 | 0.053
BS-BP | 0.1 | 0.153
BP-1 | 0.188 | 0.341
2025-01-02 | 0.371 | 0.712
2025-02-03 | 0.297 | 1.009
2025-03-04 | 0.455 | 1.464
2025-04-05 | 0.273 | 1.737
2025-05-06 | 0.215 | 1.952
2025-06-07 | 0.224 | 2.176
2025-07-08 | 0.243 | 2.419
2025-08-09 | 0.266 | 2.685
2025-09-10 | 0.312 | 2.997
2025-10-11 | 0.335 | 3.332
2025-11-12 | 0.354 | 3.686
2025-12-13 | 0.337 | 4.023
13-14 | 0.209 | 4.232
14-15 | 0.211 | 4.443
15-16 | 0.229 | 4.672
16-17 | 0.234 | 4.906
17-18 | 0.263 | 5.169
18-19 | 0.245 | 5.414
19-20 | 0.238 | 5.652
20-21 | 0.278 | 5.93
21-22 | 0.264 | 6.194
22-23 | 0.339 | 6.533
23-24 | 0.29 | 6.823
24-25 | 0.323 | 7.146
25-26 | 0.286 | 7.432
26-27 | 0.22 | 7.652
27-28 | 0.342 | 7.994
28-29 | 0.304 | 8.298
29-30 | 0.364 | 8.662
30-31 | 0.205 | 8.867
31-BB2 | 0.098 | 8.965

## Coil 2 Resistance 7_8_25
Layer | Voltage Difference | Total Voltage
BB1
BB1-BS | 0.049 | 0.049
BS-BP | 0.145 | 0.194
BP-1 | 0.319 | 0.513
1–2 | 0.375 | 0.888
2–3 | 0.239 | 1.127
3–4 | 0.321 | 1.448
4–5 | 0.278 | 1.726
5–6 | 0.178 | 1.904
6–7 | 0.195 | 2.099
7–8 | 0.193 | 2.292
8–9 | 0.217 | 2.509
9–10 | 0.25 | 2.759
10–11 | 0.332 | 3.091
11–12 | 0.325 | 3.416
12–13 | 0.334 | 3.75
13-14 | 0.184 | 3.934
14-15 | 0.281 | 4.215
15-16 | 0.192 | 4.407
16-17 | 0.176 | 4.583
17-18 | 0.254 | 4.837
18-19 | 0.295 | 5.132
19-20 | 0.267 | 5.399
20-21 | 0.251 | 5.65
21-22 | 0.235 | 5.885
22-23 | 0.268 | 6.153
23-24 | 0.318 | 6.471
24-25 | 0.467 | 6.938
25-26 | 0.212 | 7.15
26-27 | 0.232 | 7.382
27-28 | 0.347 | 7.729
28-29 | 0.349 | 8.078
29-30 | 0.478 | 8.556
30-31 | 0.318 | 8.874
31-BB2 | 0.051 | 8.925

## Coil 1 Resistance 6_25_25
Layer | Total Resistance (mΩ) | Resistance (mΩ)
BB1 | 0
Screw | 0.036 | 0.036
Plate | 0.176 | 0.14
1 | 0.379 | 0.203
2 | 0.861 | 0.482
3 | 1.112 | 0.251
4 | 1.483 | 0.371
5 | 1.801 | 0.318
6 | 1.988 | 0.187
7 | 2.171 | 0.183
8 | 2.393 | 0.222
9 | 2.515 | 0.122
10 | 2.682 | 0.167
11 | 3.034 | 0.352
12 | 3.377 | 0.343
13 | 3.733 | 0.356
14 | 3.928 | 0.195
15 | 4.268 | 0.34
16 | 4.484 | 0.216
17 | 4.663 | 0.179
18 | 4.791 | 0.128
19 | 5.102 | 0.311
20 | 5.434 | 0.332
21 | 5.714 | 0.28
22 | 6.289 | 0.575
23 | 6.493 | 0.204
24 | 6.883 | 0.39
25 | 7.2 | 0.317
26 | 7.423 | 0.223
27 | 7.71 | 0.287
28 | 8.012 | 0.302
29 | 8.293 | 0.281
30 | 8.374 | 0.081
31 | 8.579 | 0.205
BB2 | 8.656 | 0.077

## Coil 2 Resistance 6_24_25
Trial 1 |  |  | Trial 2
Layer | Total Resistance (mΩ) | Resistance (mΩ) | Layer | Total Resistance (mΩ) | Resistance (mΩ)
BB1-Sc | 0.039 |  | BB1-Sc | 0.039
Plate | 0.161 | 0.122 | Plate | 0.161 | 0.122
1 | 0.245 | 0.084 | 1 | 0.245 | 0.084
2 | 0.458 | 0.213 | 2 | 0.458 | 0.213
3 | 0.659 | 0.201 | 3 | 0.659 | 0.201
4 | 0.843 | 0.184 | 4 | 0.843 | 0.184
5 | 1.034 | 0.191 | 5 | 1.034 | 0.191
6 | 1.325 | 0.291 | 6 | 1.325 | 0.291
7 | 1.684 | 0.359 | 7 | 1.684 | 0.359
8 | 1.896 | 0.212 | 8 | 1.896 | 0.212
9 | 2.004 | 0.108 | 9 | 2.004 | 0.108
10 | 2.146 | 0.142 | 10 | 2.146 | 0.142
11 | 2.345 | 0.199 | 11 | 2.345 | 0.199
12 | 2.55 | 0.205 | 12 | 2.55 | 0.205
13 | 2.757 | 0.207 | 13 | 2.757 | 0.207
14 | 2.96 | 0.203 | 14 | 2.96 | 0.203
15 | 3.174 | 0.214 | 15 | 3.174 | 0.214
16 | 3.44 | 0.266 | 16 | 3.44 | 0.266
17 | 3.718 | 0.278 | 17 | 3.718 | 0.278
18 | 3.882 | 0.164 | 18 | 3.882 | 0.164
19 | 4.085 | 0.203 | 19 | 4.085 | 0.203
20 | 4.259 | 0.174 | 20 | 4.259 | 0.174
21 | 4.569 | 0.31 | 21 | 4.569 | 0.31
22 | 4.85 | 0.281 | 22 | 4.85 | 0.281
23 | 5.003 | 0.153 | 23 | 5.003 | 0.153
24 | 5.367 | 0.364 | 24 | 5.367 | 0.364
25 | 5.588 | 0.221 | 25 | 5.791 | 0.424
26 | 5.785 | 0.197 | 26 | 6.2 | 0.409
27 | 6.38 | 0.595 | 27 | 6.622 | 0.422
28 | 6.551 | 0.171 | 28 | 6.829 | 0.207
29 | 6.699 | 0.148 | 29 | 7.017 | 0.188
30 | 7.111 | 0.412 | 30 | 7.303 | 0.286
31 | 7.539 | 0.428 | 31 | 7.689 | 0.386
BB2 | 7.886 | 0.347 | BB2 | 7.862 | 0.173

## Coil Resistance 6_20_25
BB1-Sc | 0.201 | 0.201
Sc-BP | 0.306 | 0.105
BP-1 | 0.464 | 0.158
1–2 | 0.812 | 0.348
2–3 | 1.051 | 0.239
3–4 | 1.44 | 0.389
4–5 | 1.613 | 0.173
5–6 | 1.792 | 0.179
6–7 | 1.968 | 0.176
7–8 | 2.187 | 0.219
8–9 | 2.393 | 0.206
9–10 | 2.628 | 0.235
10–11 | 2.952 | 0.324
11–12 | 3.253 | 0.301
12–13 | 3.423 | 0.17
13-14 | 3.62 | 0.197
14-15 | 3.96 | 0.34
15-16 | 4.185 | 0.225
16-17 | 4.363 | 0.178
17-18 | 4.605 | 0.242
18-19 | 4.775 | 0.17
19-20 | 5.011 | 0.236
20-21 | 5.324 | 0.313
21-22 | 5.591 | 0.267
22-23 | 5.838 | 0.247
23-24 | 6.196 | 0.358
24-25 | 6.355 | 0.159
25-26 | 6.508 | 0.153
26-27 | 6.72 | 0.212
27-28 | 6.953 | 0.233
29-28 | 7.167 | 0.214
29-30 | 7.475 | 0.308
30-31 | 7.644 | 0.169
31-End of BB2 | 7.855 | 0.211

## Coil Resistance 4_24_25
Screw | 0.069
Plate | 0.107 | 0.038 | 1
1 | 0.428 | 0.321 | 2
2 | 0.806 | 0.378 | 3
3 | 0.912 | 0.106 | 4
4 | 1.234 | 0.322 | 5
5 | 1.478 | 0.244 | 6
6 | 1.648 | 0.17 | 7
7 | 1.825 | 0.177 | 8
8 | 2.017 | 0.192 | 9
9 | 2.223 | 0.206 | 10
10 | 2.42 | 0.197 | 11
11 | 2.547 | 0.127 | 12
12 | 2.84 | 0.293 | 13
13 | 3.145 | 0.305 | 14
14 | 3.319 | 0.174 | 15
15 | 3.602 | 0.283 | 16
16 | 3.819 | 0.217 | 17
17 | 4.032 | 0.213 | 18
18 | 4.254 | 0.222 | 19
19 | 4.558 | 0.304 | 20
20 | 4.761 | 0.203 | 21
21 | 4.98 | 0.219 | 22
22 | 5.235 | 0.255 | 23
23 | 5.43 | 0.195 | 24
24 | 5.618 | 0.188 | 25
25 | 6.038 | 0.42 | 26
26 | 6.25 | 0.212 | 27
27 | 6.405 | 0.155 | 28
28 | 6.603 | 0.198 | 29
29 | 6.883 | 0.28 | 30
30 | 7.024 | 0.141 | 31
31 | 7.22 | 0.196 | 32
BB2 | 7.373 | 0.153 | 33

## Coil Resistance 10_24_24
Layer | Voltage (mv) | Resistance (mOhms) | Layer_Resistance
BB1 St | 0.246 | 0.024724
B_Sc | 0.305 | 0.030653 | 0.0059296
BP | 1.869 | 0.18784 | 0.15719
1 | 3.958 | 0.39779 | 0.20995
2 | 6.69 | 0.67236 | 0.27457
3 | 7.707 | 0.77457 | 0.10221
4 | 9.606 | 0.96543 | 0.19085
5 | 12.062 | 1.2123 | 0.24683
6 | 14.706 | 1.478 | 0.26573
7 | 17.391 | 1.7478 | 0.26985
8 | 19.975 | 2.0075 | 0.2597
9 | 22.495 | 2.2608 | 0.25327
10 | 24.698 | 2.4822 | 0.22141
11 | 27.055 | 2.7191 | 0.23688
12 | 29.413 | 2.9561 | 0.23698
13 | 31.699 | 3.1858 | 0.22975
14 | 33.905 | 3.4075 | 0.22171
15 | 36.239 | 3.6421 | 0.23457
16 | 39.376 | 3.9574 | 0.31528
17 | 41.881 | 4.2091 | 0.25176
18 | 44.421 | 4.4644 | 0.25528
19 | 46.891 | 4.7127 | 0.24824
20 | 49.403 | 4.9651 | 0.25246
21 | 51.253 | 5.1511 | 0.18593
22 | 53.747 | 5.4017 | 0.25065
23 | 56.062 | 5.6344 | 0.23266
24 | 58.998 | 5.9294 | 0.29508
25 | 61.452 | 6.1761 | 0.24663
26 | 63.84 | 6.4161 | 0.24
27 | 65.32 | 6.5648 | 0.14874
28 | 68.033 | 6.8375 | 0.27266
29 | 70.348 | 7.0702 | 0.23266
30 | 73.303 | 7.3671 | 0.29698
31 | 75.861 | 7.6242 | 0.25709
BB2 St | 76.909 | 7.7295 | 0.10533
BB2 End | 77.076 | 7.7463 | 0.016784

## Coil Resistance 5_2_24
1 amp applied
Resistance
Total Resistance: |  | 7.053 mohms
Beginning to end of BB1 (w Cu187 screw) |  | 0.016 mohms
Beginning of BB1 to CU187 screw |  | 0.018 mohms
Top of Cu-Cr screw to Cu-Cr plate bottom |  | 0.159 mohms
Bottom to top of Cr-Cu plate |  | 0.005 mohms
Opposite ends of Cr-Cu plate |  | 0.032 mohms
Layers | Layer Resistance (mOhms) |  |  | Resistance from 4/23/24
|  |  | BB1-BP | 0.36633
Cu-Cr(bottom)-1 | 0.222 |  |  | 0.386
1–2 | 0.228 |  |  | 0.30433
2–3 | 0.254 |  |  | 0.23567
3–4 | 0.243 |  |  | 0.25467
4–5 | 0.253 |  |  | 0.25
5–6 | 0.242 |  |  | 0.25933
6–7 | 0.246 |  |  | 0.29
7–8 | 0.253 |  |  | 0.17367
8–9 | 0.235 |  |  | 0.28867
9–10 | 0.224 |  |  | 0.25867
10–11 | 0.239 |  |  | 0.24233
11–12 | 0.246 |  |  | 0.27
12–13 | 0.244 |  |  | 0.24667
13-14 | 0.247 |  |  | 0.20767
14-15 | 0.243 |  |  | 0.23433
15-16 | 0.246 |  |  | 0.25833
16-17 | 0.224 |  |  | 0.598
17-18 | 0.223 |  |  | 0.267
18-19 | 0.248 |  |  | 0.28133
19-20 | 0.242 |  |  | 0.20833
21-22 | 0.255 |  |  | 0.24967
22-23 | 0.254 |  |  | 0.284
23-24 | 0.257 |  |  | 0.122
24-25 | 0.252 |  |  | 0.346
25-26 | 0.293 |  |  | 0.18333
26-27 | 0.301 |  |  | 0.29267
27-28 | 0.258 |  |  | 0.20233
29-28 | 0.255 |  |  | 0.552
29-30 | 0.207 |  |  | 0.383
30-31 | 0.273 |  |  | 0.21667
31-End of BB2 | 0.036 |  |  | 0.25167
Total | 7.443 |  |  | 8.9647
Mean |  | 0.24776
Median: |  | 0.246
Mode |  | 0.246
Standard Deviation |  | 0.019275
Variance |  | 0.00037154

## Coil Resistance 4_24_24
Layer Resistance (mOhms)
BB1-BP | 0.36633
Cu-Cr(bottom)-1 | 0.386
1–2 | 0.30433
2–3 | 0.23567
3–4 | 0.25467
4–5 | 0.25
5–6 | 0.25933
6–7 | 0.29
7–8 | 0.17367
8–9 | 0.28867
9–10 | 0.25867
10–11 | 0.24233
11–12 | 0.27
12–13 | 0.24667
13-14 | 0.20767
14-15 | 0.23433
15-16 | 0.25833
16-17 | 0.598
17-18 | 0.267
18-19 | 0.28133
19-20 | 0.20833
21-22 | 0.24967
22-23 | 0.284
23-24 | 0.122
24-25 | 0.346
25-26 | 0.18333
26-27 | 0.29267
27-28 | 0.20233
29-28 | 0.552
29-30 | 0.383
30-31 | 0.21667
31-End of BB2 | 0.25167
Total | 8.9647

## Coil Resistance 5_2_23
3 amps applied
Layer | Volt Difference from Ground (mV) | Resistance from Previous Layer (mOhms) |  |  |  | Uncertainty (from measuring brass plate 5 times)
Ground | 0.015 |  |  |  |  | Brass Plate Measurements
Bus Bar 1 | 66.459 | 22.148 | Average | 0.2665 |  | 1 | 67.64
Brass Screw | 66.511 | 0.017333 | Maximum: | 0.598 |  | 2 | 67.45
Brass Plate | 67.61 | 0.36633 | Max % of mean | 224.3902 |  | 3 | 67.772
1 | 68.768 | 0.386 | Minimum | 0.122 |  | 4 | 67.482
2 | 69.681 | 0.30433 | Min % of Mean | 0.45779 |  | 5 | 67.625
3 | 70.388 | 0.23567 |  |  |  | Average | 67.5938
4 | 71.152 | 0.25467 | Total Coil Resistance |  | 9.061 |  | 0.00010672
5 | 71.902 | 0.25 |  |  |  |  | 0.0010339
6 | 72.68 | 0.25933 | Spikes |  |  |  | 0.0015878
7 | 73.55 | 0.29 | Layer 17 |  |  |  | 0.00062496
8 | 74.071 | 0.17367 | Layer 28 |  |  |  | 4.8672e-05
9 | 74.937 | 0.28867
10 | 75.713 | 0.25867 |  |  |  | Uncertainty | 0.058327
11 | 76.44 | 0.24233
12 | 77.25 | 0.27
13 | 77.99 | 0.24667
14 | 78.613 | 0.20767
15 | 79.316 | 0.23433
16 | 80.091 | 0.25833
17 | 81.885 | 0.598
18 | 82.686 | 0.267
19 | 83.53 | 0.28133
20 | 84.155 | 0.20833
21 | 84.904 | 0.24967
22 | 85.756 | 0.284
23 | 86.122 | 0.122
24 | 87.16 | 0.346
25 | 87.71 | 0.18333
26 | 88.588 | 0.29267
27 | 89.195 | 0.20233 | 2.0713
28 | 90.851 | 0.552
29 | 92 | 0.383
30 | 92.65 | 0.21667
31 | 93.405 | 0.25167
Bus Bar 2 | 93.642 | 0.079
Positive Terminal | 130.342 | 12.2333

## Coil Resistance 4_26_23
Current Applied | 3 amps
Ground voltage: | 0.002 mV
Positive Terminal: | 141.90 mV |  |  |  | 1432
Input Bus Bar (mV): | 64.475 | Voltage Difference of Coil; | 26.85 | mV
Output Bus Bar (mV) | 91.325
|  | Resistance of Coil | 0.00895 | Ohms
Coil Layer Measurements (mV) |  | Predicted Heat with 400 Amps | 1432 | Watts |  |  | 0.087745
Input Bus Bar (mV): | 64.475 | Voltage diff (mV) | Resistance (mOhms) |  | Average Resistance | 0.26324 | mV
Brass Screw (mV) | 64.618 | 0.143 | 0.047667 |  | Maximum Resistance: | 0.45667 | mV at 29-30 plates
Brass Plate (mV) | 65.635 | 1.017 | 0.339 |  | Deviation of max: | 173.4823 | % | of Average
1 | 66.571 | 0.936 | 0.312 |  | Minimum Resistance: | 0.032667
2 | 67.335 | 0.764 | 0.25467 |  | Deviation of Min | 7.1533 | % | of Average | at 7-8 plates
3 | 68.34 | 1.005 | 0.335
4 | 69.191 | 0.851 | 0.28367
5 | 70.082 | 0.891 | 0.297
6 | 70.83 | 0.748 | 0.24933
7 | 71.75 | 0.92 | 0.30667
8 | 71.848 | 0.098 | 0.032667 | 0.1241
9 | 72.55 | 0.702 | 0.234
10 | 73.355 | 0.805 | 0.26833
11 | 74.06 | 0.705 | 0.235
12 | 74.264 | 0.204 | 0.068 |  | 0.25832
13 | 75.051 | 0.787 | 0.26233
14 | 75.729 | 0.678 | 0.226
15 | 76.512 | 0.783 | 0.261
16 | 77.305 | 0.793 | 0.26433
17 | 77.851 | 0.546 | 0.182
18 | 78.275 | 0.424 | 0.14133
19 | 79.182 | 0.907 | 0.30233
20 | 80.207 | 1.025 | 0.34167
21 | 80.902 | 0.695 | 0.23167
22 | 81.84 | 0.938 | 0.31267
23 | 82.82 | 0.98 | 0.32667
24 | 83.6 | 0.78 | 0.26
25 | 84.34 | 0.74 | 0.24667
26 | 84.977 | 0.637 | 0.21233
27 | 86.033 | 1.056 | 0.352
28 | 87.277 | 1.244 | 0.41467
29 | 88.33 | 1.053 | 0.351
30 | 89.7 | 1.37 | 0.45667
31 | 90.67 | 0.97 | 0.32333
Bus Bar 2 | 91.325 | 0.655 | 0.21833
Uncertainty: Measured from multiple brass plate measurements |  |  | 1 | 66.14 | mV
|  |  | 2 | 66.368 | mV
|  |  | 3 | 66.3 | mV
|  |  | 4 | 66.295 | mV
|  |  | 5 | 66.275 | mV
|  |  | Max | 66.368 | mV
|  |  | Min | 66.14 | mV
|  |  | Uncertainty | 0.114 | mV

## Temps 5_2_24 (WW)
Current | Power | Voltage | Plate | Screw | Coil | Tube | Bus Bar
0 | 0 | 0 | 14.9 | 19.9 | 15 | 14.6 | 19.7 | Wall Water and No Bus Bar Cooling | (~45 PSI, 2.0 GPM)
100 | 71 | 0.722 | 15.1 | 20 | 15.3 | 15.1 | 19.8
135 | 129.3975 | 0.974 | 15.2 | 20.2 | 15.9 | 15 | 19.7
160 | 181.76 | 1.161 | 15.7 | 20.6 | 16.1 | 15.1 | 20.6
175 | 217.4375 | 1.269 | 16.4 | 20.6 | 16.1 | 15.4 | 21.3
190 | 256.31 | 1.382 | 16.5 | 21.2 | 16.3 | 15.4 | 22.1
200 | 284 | 1.456 | 16.7 | 21.3 | 16.4 | 15.8 | 22.9
210 | 313.11 | 1.53 | 16.7 | 21.6 | 16.5 | 16.6 | 23.3
220 | 343.64 | 1.604 | 16.8 | 22.1 | 16.7 | 16.6 | 23.6
230 | 375.59 | 1.677 | 17.1 | 22.7 | 16.9 | 16.6 | 24.2
240 | 408.96 | 1.751 | 17.3 | 23.4 | 17.2 | 16.7 | 25.3
250 | 443.75 | 1.829 | 17.7 | 24.3 | 17.5 | 17.2 | 26.1
260 | 479.96 | 1.899 | 17.7 | 24.7 | 17.5 | 17.2 | 26.4
270 | 517.59 | 1.977 | 17.8 | 25.5 | 17.7 | 17.3 | 27.4
280 | 556.64 | 2.055 | 18.1 | 26.4 | 17.8 | 17.3 | 28.4
290 | 597.11 | 2.129 | 18.2 | 27.3 | 18 | 17.3 | 29.2
300 | 639 | 2.207 | 18.3 | 28.4 | 18.1 | 17.3 | 30.5
310 | 682.31 | 2.285 | 18.6 | 29.1 | 18.4 | 17.4 | 31.2
320 | 727.04 | 2.359 | 18.6 | 29.7 | 18.6 | 17.4 | 31.9
330 | 773.19 | 2.437 | 18.9 | 30.6 | 18.9 | 17.4 | 32.8
Temperature Change © |  |  | 4 | 10.7 | 3.9 | 2.8 | 13.1

## Temps 5_2_24 (2) (WW)
Current | Power | Voltage | Bus Bar | Water | Coil | Plate | Screw
0 | 0 | 0 | 21.3 | 16.2 | 16.2 | 15.8 | 21.5
49.99 | 19.992 | 0.388 | 21.2 | 17 | 16.9 | 16.3 | 21.2
75 | 45 | 0.579 | 21.3 | 17.1 | 17 | 17.1 | 21.3
100.01 | 80.016 | 0.774 | 22.3 | 17.5 | 17.6 | 18.3 | 22.1
110.11 | 96.9937 | 0.853 | 23.3 | 17.8 | 18 | 19 | 23
119.97 | 115.1424 | 0.935 | 23.8 | 17.9 | 18.1 | 19.6 | 23.4
129.95 | 135.096 | 1.009 | 24.3 | 17.7 | 18.2 | 20.2 | 23.9
139.96 | 156.7104 | 1.083 | 25.1 | 18.3 | 18.6 | 21 | 24.4
150.03 | 180.072 | 1.179 | 26.5 | 18.6 | 18.9 | 22 | 25.6
160 | 204.8 | 1.252 | 27.1 | 18.6 | 19.2 | 22.7 | 26.2
169.98 | 231.1456 | 1.334 | 27.8 | 18.7 | 19.5 | 23.7 | 26.7
179.96 | 259.0848 | 1.417 | 28.8 | 19.2 | 20 | 24.7 | 27.8
190.06 | 288.9824 | 1.499 | 30.3 | 19.7 | 20.4 | 25.8 | 29.1
200.04 | 320.128 | 1.582 | 31.4 | 19.5 | 20.7 | 26.7 | 30.1
Temperature Differences |  |  | 10.1 | 3.3 | 4.5 | 10.9 | 8.6

## Temps 4_26_24 (WW)
Current | Power | Voltage | Cu-Cr Plate | Cu-Cr Screw | Coil Body (lower half) | Tube | BB2 | BB1 (con to Cu-Cr screw)
Trial 1
0 | 0 |  | 21.2 | 21 | 21 | 22.2 | 21.1 | 21.4 | Cooling off |  | (~45 PSI, 2.0 GPM)
0 | 0 | 0 | 15.1 | 19.4 | 14.6 | 14.7 | 18.7 |  | Cooling on | Wall Water and No Bus Bar Cooling
100 | 71 | 0.722 | 15.2 | 20 | 15.2 | 14.8 | 19.2
135 | 129.3975 | 0.974 | 15.5 | 20.1 | 15.7 | 15 | 19.9
160 | 181.76 | 1.161 | 15.9 | 20.3 | 15.9 | 15 | 20..6
175 | 217.4375
190 | 256.31
200 | 284 |  | 16.7 | 20.9 | 16 | 15.5 | 21.9
Trial 2
0 | 0 | 0 | 14.9 | 19.9 | 15 | 14.6 | 19.7
100 | 71 | 0.722 | 15.1 | 20 | 15.3 | 15.1 | 19.8
135 | 129.3975 | 0.974 | 15.2 | 20.2 | 15.9 | 15 | 19.7
160 | 181.76 | 1.161 | 15.7 | 20.6 | 16.1 | 15.1 | 20.6
175 | 217.4375 | 1.269 | 16.4 | 20.6 | 16.1 | 15.4 | 21.3
190 | 256.31 | 1.382 | 16.5 | 21.2 | 16.3 | 15.4 | 22.1
200 | 284 | 1.456 | 16.7 | 21.3 | 16.4 | 15.8 | 22.9
210 | 313.11 | 1.53 | 16.7 | 21.6 | 16.5 | 16.6 | 23.3
220 | 343.64 | 1.604 | 16.8 | 22.1 | 16.7 | 16.6 | 23.6
230 | 375.59 | 1.677 | 17.1 | 22.7 | 16.9 | 16.6 | 24.2
240 | 408.96 | 1.751 | 17.3 | 23.4 | 17.2 | 16.7 | 25.3
250 | 443.75 | 1.829 | 17.7 | 24.3 | 17.5 | 17.2 | 26.1
260 | 479.96 | 1.899 | 17.7 | 24.7 | 17.5 | 17.2 | 26.4
270 | 517.59 | 1.977 | 17.8 | 25.5 | 17.7 | 17.3 | 27.4
280 | 556.64 | 2.055 | 18.1 | 26.4 | 17.8 | 17.3 | 28.4
290 | 597.11 | 2.129 | 18.2 | 27.3 | 18 | 17.3 | 29.2
300 | 639 | 2.207 | 18.3 | 28.4 | 18.1 | 17.3 | 30.5
310 | 682.31 | 2.285 | 18.6 | 29.1 | 18.4 | 17.4 | 31.2
320 | 727.04 | 2.359 | 18.6 | 29.7 | 18.6 | 17.4 | 31.9
330 | 773.19 | 2.437 | 18.9 | 30.6 | 18.9 | 17.4 | 32.8
Temperature Change © |  |  | 4 | 10.7 | 3.9 | 2.8 | 13.1

## Temps 1_25_24 (4) (WW)
Current | Cu-Cr Plate | Cu-Cr Screw | Coil Body (lower half) | Tube | BB2 | BB1 (con to Cu-Cr screw)
0 | 20.7 | 16.2 | 16.2 | 15.6 | 21 | 22 | (~45 PSI, 2.0 GPM)
49.99 | 20.8 | 16.5 | 16.4 | 15.7 | 21.2 | 22.2 |  | Wall Water and No Bus Bar Cooling
100.01 | 21 | 17.2 | 17.3 | 17.6 | 21.6 | 22.6
115.04 | 21.4 | 17.5 | 17.5 | 18.8 | 21.9 | 23.1
130.07 | 21.9 | 17.7 | 18.3 | 19.7 | 22.2 | 23.1
145.1 | 22.1 | 17.9 | 18.7 | 20.7 | 22.5 | 23.2
160 | 23.3 | 18.3 | 19.1 | 22.3 | 23.6 | 24.4
175.03 | 24.3 | 18.7 | 19.8 | 23.7 | 24.4 | 25.1
190.06 | 25 | 19.3 | 20.5 | 25.2 | 25.2 | 25.9
200.04 | 26 | 19.6 | 21 | 26.5 | 26.2 | 26.8
Temp Ch | 5.3 | 3.4 | 4.8 | 10.9 | 5.2 | 4.8

## Temps 1_25_24 (3) (WW)
Current | Cu-Cr Plate | Cu-Cr Screw | Coil Body (lower half) | Tube | BB2 | BB1 (con to Cu-Cr screw)
0 | 21.3 | 16.2 | 16.2 | 15.8 | 21.5 | 22.1
49.99 | 21.2 | 17 | 16.9 | 16.3 | 21.2 | 21.8
75 | 21.3 | 17.1 | 17 | 17.1 | 21.3 | 21.9
100.01 | 22.3 | 17.5 | 17.6 | 18.3 | 22.1 | 22.9
110.11 | 23.3 | 17.8 | 18 | 19 | 23 | 23.8
119.97 | 23.8 | 17.9 | 18.1 | 19.6 | 23.4 | 24
129.95 | 24.3 | 17.7 | 18.2 | 20.2 | 23.9 | 24.5
139.96 | 25.1 | 18.3 | 18.6 | 21 | 24.4 | 25.2
150.03 | 26.5 | 18.6 | 18.9 | 22 | 25.6 | 26.3
160 | 27.1 | 18.6 | 19.2 | 22.7 | 26.2 | 26.4
169.98 | 27.8 | 18.7 | 19.5 | 23.7 | 26.7 | 27.2
179.96 | 28.8 | 19.2 | 20 | 24.7 | 27.8 | 28.7
190.06 | 30.3 | 19.7 | 20.4 | 25.8 | 29.1 | 29.1
200.04 | 31.4 | 19.5 | 20.7 | 26.7 | 30.1 | 30.3

## B-Field Coil 2 8_04_25 New Coil
Distance from Top (cm) | Gauss | Imperial | Distance from bottom of coil |  |  | 100.01 A, 0.801 V applied to Coil 2
0 | 120.67 | 0 | -74.8538 |  | Top of plastic piece
1 | 177.93 | 0.3937 | -64.8538
2 | 205.44 | 0.7874 | -54.8538
3 | 262.84 | 1.1811 | -44.8538 |  |  | Newest version of coil constructed, but measurement error due to a misaligned probe
3.5 | 288.71 | 1.378 | -39.8538
4 | 299.64 | 1.5748 | -34.8538
4.25 | 330.52 | 1.6732 | -32.3538
4.5 | 340.53 | 1.7717 | -29.8538
4.75 | 345.44 | 1.8701 | -27.3538
5 | 358.05 | 1.9685 | -24.8538
5.25 | 361.51 | 2.0669 | -22.3538 |  | Peak field 22.3538 mm from Coil Bottom
5.5 | 366.06 | 2.1654 | -19.8538
5.75 | 365.15 | 2.2638 | -17.3538
6 | 359.86 | 2.3622 | -14.8538
6.25 | 352.54 | 2.4606 | -12.3538
6.5 | 341.51 | 2.5591 | -9.8538
6.75 | 328.94 | 2.6575 | -7.3538
7 | 315.36 | 2.7559 | -4.8538
7.25 | 294.08 | 2.8543 | -2.3538
7.5 | 284.08 | 2.9528 | 0.1462 | Approximate Coil Bottom location (Coil is 2.947 inches tall)
7.62 | 278.67 | 3 | 1.3462
7.7787 | 265.16 | 3.0625 | 2.9337
7.9375 | 248.86 | 3.125 | 4.5212
8.0962 | 234.76 | 3.1875 | 6.1087
8.255 | 218.9 | 3.25 | 7.6962
8.4138 | 207.59 | 3.3125 | 9.2837
8.5725 | 196.97 | 3.375 | 10.8712
8.7186 | 181.54 | 3.4325 | 12.3317
8.89 | 171.89 | 3.5 | 14.0462
9.0488 | 163.26 | 3.5625 | 15.6337
9.2075 | 156.28 | 3.625 | 17.2212 | Approximate Atom location (17 mm from coil bottom)
9.3663 | 144.78 | 3.6875 | 18.8087
9.525 | 130.69 | 3.75 | 20.3962
9.6837 | 121.02 | 3.8125 | 21.9837
9.8425 | 114.63 | 3.875 | 23.5712
10.0013 | 107.52 | 3.9375 | 25.1587
10.16 | 101.71 | 4 | 26.7462
10.795 | 78.15 | 4.25 | 33.0962
11.43 | 63.3 | 4.5 | 39.4462
12.065 | 49.97 | 4.75 | 45.7962
12.7 | 43.46 | 5 | 52.1462
13.97 | 31.39 | 5.5 | 64.8462
15.24 | 24.8 | 6 | 77.5462

## Coil B_Field 5_9_24
Distance from Coil Bottom (in) | B-Field (G) | Distance from Coil Bottom (cm)
-4 | 58.75 | -10.16
-3.8 | 71.07 | -9.652
-3.6 | 82.59 | -9.144
-3.4 | 107.13 | -8.636
-3.2 | 126.49 | -8.128
-3 | 149.98 | -7.62
-2.8 | 180.38 | -7.112
-2.6 | 216.75 | -6.604
-2.4 | 253.52 | -6.096
-2.2 | 291.6 | -5.588
-2 | 315.28 | -5.08
-1.8 | 334.22 | -4.572
-1.6 | 373.93 | -4.064
-1.4 | 386.97 | -3.556
-1.2 | 392.41 | -3.048 | Peak Field -3.048 cm from coil Bottom
-1 | 389.4 | -2.54
-0.8 | 387.04 | -2.032
-0.6 | 375.24 | -1.524
-0.4 | 348.65 | -1.016
-0.2 | 313.45 | -0.508
0 | 277.63 | 0 | Coil Bottom Field
0.2 | 234.8 | 0.508
0.4 | 194.46 | 1.016
0.6 | 170.17 | 1.524 | Approximate field at atom location (17 mm)
0.8 | 137.9 | 2.032
1 | 113.36 | 2.54
1.2 | 94.52 | 3.048
1.4 | 78.15 | 3.556
1.6 | 65.36 | 4.064
1.8 | 55.81 | 4.572
2 | 46.69 | 5.08

## Coil B_Field 11_27_23
Distance from Coil Bottom (cm) | B-Field (G) | B-Field Error (G)
-7.62 | 101.39 | 5.2
-7.112 | 120.11 | 5.2
-6.604 | 150.56 | 5.2
-6.096 | 187.48 | 5.2
-5.588 | 217.86 | 5.2
-5.08 | 265.62 | 5.2
-4.572 | 318.43 | 5.2
-4.064 | 358.82 | 5.2
-3.556 | 391.7 | 5.2
-3.048 | 419.13 | 5.2
-2.54 | 423.01 | 5.2
-2.032 | 409.51 | 5.2
-1.524 | 383.34 | 5.2
-1.016 | 345.96 | 5.2
-0.508 | 309 | 5.2
0 | 271.08 | 5.2
0.2 | 257.11 | 5.2
0.4 | 241.73 | 5.2
0.6 | 222.58 | 5.2
0.8 | 210.4 | 5.2
1 | 195.9 | 5.2
1.2 | 186 | 5.2
1.4 | 172.13 | 5.2
1.5 | 161.33 | 5.2
1.6 | 154.26 | 5.2
1.8 | 145.06 | 5.2
2 | 129.94 | 5.2
2.2 | 126.54 | 5.2
2.4 | 119.81 | 5.2
2.6 | 111.23 | 5.2
2.8 | 104.12 | 5.2
3 | 97.01 | 5.2

## Coil B_Fields 11_27_23 (2)
Distance from Coil Bottom (cm) | B-Field (G) |  | Distance from Coil Bottom (cm) | B-Field (G) |  | Distance from Coil Bottom (cm) | B-Field (G) |  | Distance from Coil Bottom (cm) | B-Field (G)
-7.112 | 131.98 |  | -7.62 | 102.95 |  | -7.62 | 101.39 |  | -7.62 | 102.17
-6.604 | 159.96 |  | -7.112 | 126.84 |  | -7.112 | 120.11 |  | -7.112 | 126.31
-6.096 | 195.91 |  | -6.604 | 148.27 |  | -6.604 | 150.56 |  | -6.604 | 152.93
-5.588 | 230.79 |  | -6.096 | 184.74 |  | -6.096 | 187.48 |  | -6.096 | 189.3767
-5.08 | 271.73 |  | -5.588 | 221.78 |  | -5.588 | 217.86 |  | -5.588 | 223.4767
-4.572 | 317.77 |  | -5.08 | 265.63 |  | -5.08 | 265.62 |  | -5.08 | 267.66
-4.064 | 362 |  | -4.572 | 307.34 |  | -4.572 | 318.43 |  | -4.572 | 314.5133
-3.556 | 396.21 |  | -4.064 | 352.26 |  | -4.064 | 358.82 |  | -4.064 | 357.6933
-3.048 | 408.35 |  | -3.556 | 386.08 |  | -3.556 | 391.7 |  | -3.556 | 391.33
-2.54 | 419.24 | Peak Field | -3.048 | 416.4 |  | -3.048 | 419.13 |  | -3.048 | 414.6267
-2.032 | 406.75 |  | -2.54 | 423.35 | Peak Field | -2.54 | 423.01 | Peak Field | -2.54 | 421.8667 | Peak Field
-1.524 | 371.93 |  | -2.032 | 412.88 |  | -2.032 | 409.51 |  | -2.032 | 409.7133
-1.016 | 342.8 |  | -1.524 | 394.23 |  | -1.524 | 383.34 |  | -1.524 | 383.1667
-0.508 | 308.83 |  | -1.016 | 358.15 |  | -1.016 | 345.96 |  | -1.016 | 348.97
0 | 272.15 | Bottom | -0.508 | 313.09 |  | -0.508 | 309 |  | -0.508 | 310.3067
0.2 | 256.01 |  | 0 | 270.19 | Bottom | 0 | 271.08 |  | 0 | 271.14 | Bottom
0.4 | 245.93 |  | 0.2 | 252.24 |  | 0.2 | 257.11 |  | 0.2 | 255.12
0.6 | 228.95 |  | 0.4 | 237.01 |  | 0.4 | 241.73 |  | 0.4 | 241.5567
0.8 | 210.01 |  | 0.6 | 220.62 |  | 0.6 | 222.58 |  | 0.6 | 224.05
1 | 196.77 |  | 0.8 | 202.88 |  | 0.8 | 210.4 |  | 0.8 | 207.7633
1.2 | 181.57 |  | 1 | 188.86 |  | 1 | 195.9 |  | 1 | 193.8433
1.4 | 169.32 |  | 1.2 | 175.43 |  | 1.2 | 186 |  | 1.2 | 181
1.5 | 160.01 |  | 1.4 | 161.96 |  | 1.4 | 172.13 |  | 1.4 | 167.8033
1.6 | 147.7 | Approx | 1.5 | 156.87 |  | 1.5 | 161.33 |  | 1.5 | 157.0467
1.8 | 139.94 | Atom Loc | 1.6 | 149.49 | Approx | 1.6 | 154.26 | Approx | 1.6 | 152.84 | Approx
2 | 131.78 |  | 1.8 | 143.66 | Atom Loc | 1.8 | 145.06 | Atom Loc | 1.8 | 142.8867 | Atom Loc
|  |  | 2 | 135.3 |  | 2 | 129.94 |  | 2 | 132.34
Atom Field at 17 mm |  | 143.82 | 2.2 | 126.54 |  |  |  |  | 2.2 | 126.54
|  |  | 2.4 | 119.81 |  | Atom Field at 17 mm |  | 149.66 | 2.4 | 119.81
|  |  | 2.6 | 111.23 |  |  |  |  | 2.6 | 111.23
|  |  | 2.8 | 104.12 |  |  |  |  | 2.8 | 104.12
|  |  | 3 | 97.01 |  |  |  |  | 3 | 97.01
[9 empty] | Atom Field at 17 mm |  | 147.8633
|  |  | Atom Field at 17 mm |  | 146.575

## Coil B_Fields 11_27_23 (3)
Distance from Coil Bottom (cm) | B-Field (G)
-7.62 | 102.95
-7.112 | 126.84
-6.604 | 148.27
-6.096 | 184.74
-5.588 | 221.78
-5.08 | 265.63
-4.572 | 307.34
-4.064 | 352.26
-3.556 | 386.08
-3.048 | 416.4
-2.54 | 423.35 | Peak Field
-2.032 | 412.88
-1.524 | 394.23
-1.016 | 358.15
-0.508 | 313.09
0 | 270.19 | Coil Bottom
0.2 | 252.24
0.4 | 237.01
0.6 | 220.62
0.8 | 202.88
1 | 188.86
1.2 | 175.43
1.4 | 161.96
1.6 | 156.87 | Approximate Atom Location (17 mm)
1.8 | 149.49
2 | 143.66
2.2 | 126.54
2.4 | 119.81
2.6 | 111.23
2.8 | 104.12
3 | 97.01
Estimated B-Field at Atoms: |  |  | 153.18

## Coil B_Fields 11_27_23 (4)
Distance from Bottom (cm) | Field Strength (Gauss)
-5.4 | 262.27
-5.2 | 279.12
-5 | 293.04
-4.8 | 308.41
-4.6 | 330.52
-4.4 | 353.36
-4.2 | 364.52
-4 | 374
-3.8 | 385.79
-3.6 | 399.37
-3.4 | 406.53
-3.2 | 412.55
-3 | 417.3
-2.8 | 418.09 | Peak Field
-2.6 | 415.17
-2.4 | 415.66
-2.2 | 414.67
-2 | 413.83
-1.8 | 405.19
-1.6 | 395.16
-1.4 | 370.83
-1.2 | 358.34
-1 | 344.8
-0.8 | 333.25
-0.6 | 328.16
-0.4 | 311.25
-0.2 | 283.74
0 | 266.09 | Coil Bottom
0.2 | 253.79
0.4 | 242.76
0.6 | 214.4
0.8 | 193.09
1 | 185.6
1.2 | 120.25
1.4 | 160.37
1.6 | 148.03 | Approximate Atom Location (17 mm)
1.8 | 136.8
2 | 128.42
2.2 | 119.84
2.4 | 102.52
2.6 | 100.05
2.8 | 90.83
3 | 88.07
Average Field at Atoms (17 mm): |  |  | 142.415

## Coil B_Fields 11_6_23
x (cm from Coil Bottom) | B-Field (Gauss) |  | Ambient Field | 0.97 G
-2 | 409.7 | Peak Field | 100.01 A applied | 0.770 volts applied
-1.8 | 406.42
-1.6 | 399.74
-1.4 | 386.3
-1.2 | 379.91
-1 | 365.83
-0.8 | 344.19
-0.6 | 323.59
-0.4 | 315.54
-0.2 | 301.5
0 | 291.13 | Coil Bottom
0.2 | 266.08
0.4 | 251.63
0.6 | 246.73
0.8 | 229.34
1 | 208.09
1.2 | 184.71
1.4 | 164.92
1.6 | 146.9
1.8 | 136.47
2 | 118.81
Average Field at Atoms (17 mm) |  |  | 141.685

## Coil B_Fields 11_1_23
x (cm from Coil Bottom) | B-Field (Gauss)
-2 | 399.66 | Peak Field
-1.5 | 375.76
-1 | 342.24
-0.5 | 329.23
0 | 268.92 | Coil Bottom
0.5 | 230.48
1 | 191.93
1.5 | 158.65 | Approximate Coil Bottom Location (1.7 cm from bottom)
2 | 134.33
2.5 | 115.65
3 | 96.82
3.5 | 80.19
4 | 67.45
4.5 | 58.8
5 | 50.05
5.5 | 43.51
6 | 37.6
6.5 | 32.42
7 | 27.54
7.5 | 24.61
8 | 21.25
8.5 | 18.89
9 | 17.1
9.5 | 15.11
10 | 13.86
10.5 | 12.35
11 | 11.28
11.5 | 10.53
12 | 9.39
12.5 | 8.65
13 | 8.03
13.5 | 7.15
14 | 6.73
14.5 | 6.22
15 | 5.73
15.5 | 5.2
16 | 4.93
16.5 | 4.51
17 | 4.36
17.5 | 4.09
18 | 3.95
18.5 | 3.77

## Coil B_Fields 11_1_23 (2)
x (cm from Coil Bottom) | B-Field (Gauss)
-2 | 330.54 | Peak Field
-1 | 311.42
0 | 261.85 | Coil Bottom
0.5 | 235.38
1 | 192.27
1.5 | 153.66 | Approximate Location of Atoms (1.7 cm from coil bottom)
2 | 130.29
2.5 | 108.77
3 | 85.34
3.5 | 71.3
4 | 56.1
4.5 | 49.42
5 | 41.69
5.5 | 37.14
6 | 30.3
6.5 | 26.75
7 | 23.8
7.5 | 21.13
8 | 18.75
8.5 | 17.03
9 | 16.13
9.5 | 14.83
10 | 13.76
10.5 | 12.54
11 | 11.6
11.5 | 10.84
12 | 10.15
12.5 | 9.66
13 | 8.93
13.5 | 8.4
14 | 7.79
14.5 | 7.42
15 | 7.04
15.5 | 6.56
16 | 6.4
16.5 | 6.33
17 | 5.86
17.5 | 5.7
18 | 5.42
18.5 | 5.23

## Coil Inductance 1_25_24 (2)
Frequency | Coil Voltage | 100 Ohm Resistor Voltage | Impedence (mOhms) | Angular Frequency
2 | 4.08 | 43.2 | 9.444444444 | 12.56637061
5 | 5.66 | 56.2 | 10.07117438 | 31.41592654
10 | 6 | 66.4 | 9.036144578 | 62.83185307
20 | 7 | 72 | 9.722222222 | 125.6637061
40 | 7.4 | 82 | 9.024390244 | 251.3274123
80 | 12.6 | 84 | 15 | 502.6548246
100 | 14.8 | 78 | 18.97435897 | 628.3185307
200 | 30.2 | 74 |  | 1256.637061
400 | 72.8 | 68 | 40.81081081 | 2513.274123
800 | 152 | 64 | 107.0588235 | 5026.548246
1000 | 188 | 64 | 237.5 | 6283.185307
2000 | 356 | 56 | 293.75 | 12566.37061
4000 | 440 | 52 | 635.7142857 | 25132.74123
846.1538462